Activity 2: Calibrating equipment

This activity requires you to complete a check on the calibration of a micropipette, as an example of checking high resolution laboratory equipment for accuracy.

The practical requires you to work accurately and with precision over a number of readings, using equipment with which you may not previously be familiar.

Equipment

- P1000 micropipette

- Pipette tips

- Distilled water (50cm3)

- A beaker

- Digital thermometer or temperature probe

- Digital balance, resolution ±0.0001g

- Balance draught shield

- Weighing boats

Safety

Risk of breakages and spillages

Procedure

- Collect approximately 50cm3 distilled water in a beaker. Stir, and leave for a few minutes to ensure it reaches a consistent temperature.

- Use the digital thermometer/temperature probe to measure the temperature of the distilled water in several positions; confirm that the temperature throughout the liquid is consistent to within ±0.1°C.

- Ensure the draught shield is in place on the balance. Place an empty weighing boat onto the balance and tare to 0.0000g.

- Collect a P1000 micropipette. Select a volume of 500𝜇l.

- Attach a new tip to the micropipette.

- Fill the micropipette up to the required volume with distilled water; ensure that air bubbles have not formed in the micropipette. If bubbles are seen, fully expel the liquid and repeat this step.

- Expel the liquid from the micropipette into the weighing boat. Take care to ensure all liquid is expelled from the micropipette.

- Measure the mass of water added to the weighing boat.

- Repeat the whole process until you have five results. Discard any anomalies, then calculate the mean mass of water collected in the weighing boat.

- Using the table below, identify the correct Z-factor for your water temperature. This is a correction factor to allow for the change in water density at different temperatures.

| Temp/°C | Z-factor | Temp/°C | Z-factor | Temp/°C | Z-factor |
| --- | --- | --- | --- | --- | --- |
| 15 | 1.0019 | 19 | 1.0027 | 23 | 1.0035 |
| 16 | 1.0020 | 20 | 1.0029 | 24 | 1.0037 |
| 17 | 1.0023 | 21 | 1.0031 | 25 | 1.0039 |
| 18 | 1.0025 | 22 | 1.0033 |  |  |

- Calculate the measured volume of water from the micropipette using the formula:

measured volume (𝜇l) = mean mass of water  Z-factor  1000

- Calculate the measurement error in your result using the formula:

measurement error = measured volume (𝜇l) – 500

- Calculate the % measurement error, using the formula:

% measurement error = (measurement error / 500)  100

To be correctly calibrated, the % measurement error should lie in the range ±1%.

- Review these steps. Consider where errors may have occurred, and the possible implications of these errors. Agree your calibration findings with your teacher/facilitator.